## Slide 1
Importance of Reference Material for Assaying Precious Metals as per standards in vogue
Pankaj Deshmukh 12th Aug 2017

## Slide 2
Objectives
An MKS PAMP Group Company

## Slide 3
Objectives :
An MKS PAMP Group Company
What is Reference material
Various Standards in vogue and specifications for Reference materials
Importance of traceability
Comparison for the specifications for Gold-Silver RMs
What we can provide

## Slide 4
An MKS PAMP Group Company
Definition of Reference Material

## Slide 5
What is Reference Material :
An MKS PAMP Group Company
REFERENCE MATERIAL :
A Material or substance whose property values are sufficiently homogeneous and well established for use for calibration of an apparatus, the assessment of a measurement method, or for assigning values to materials
CERTIFIED REFERENCE MATERIALS (CRMs):
CRM is a reference material issued by a competent authority, accompanied by a certificate that provides the value of the specified property, its associated uncertainty, and a statement of metrological traceability.
STANDARD REFERENCE MATERIALS (SRM):
SRM is synonymous with reference material (note: in the past this acronym has also been used for “Secondary Reference Materials” – materials characterized to a lower level than and generally not conforming to all ISO criteria for a CRM).
QUALITY CONTROL MATERIAL (QCM):
is a material intended for internal laboratory quality control. A material or substance, one or more of whose property values are sufficiently homogeneous, stable and well established to be used for maintaining or monitoring measurement processes. A QCM does not have formally assigned property values or uncertainties (ISO Guide 80 - Guidance for in-house Production of Reference Materials for Metrological Quality Control (QCMs)).

## Slide 6
Difference between RM and CRM
An MKS PAMP Group Company
Reference Material (RM)
material, sufﬁciently homogeneous and stable with respect to one or more speciﬁed properties, which has been established to be ﬁt for its intended use in a measurement process
NOTE 1 - RM is a generic term.
NOTE 2 - Properties can be quantitative or qualitative.
NOTE 3 - Uses may include the calibration of a measurement system, assessment of a measurement procedure, assigning values to other materials, and quality control.
NOTE 4 - RM can only be used for a single purpose in a given measurement. For instance, the same RM cannot be used for both calibration and validation of results in the same measurement procedure.
NOTE 5 - VIM has an analogous definition (ISO/IEC Guide 99:2007, 5.13), but restricts the term “measurement” to apply to quantitative values and not to qualitative properties. However, Note 3 of ISO/IEC Guide 99:2007, 5.13 specifically includes the concept of qualitative attributes, called “nominal properties”.
Certiﬁed Reference Material (CRM)
reference material characterized by a metrologically valid procedure for one or more speciﬁed properties, accompanied by a certiﬁcate that provides the value of the speciﬁed property, its associated uncertainty, and a statement of metrological traceability
NOTE 1 - The concept of value includes qualitative attributes such as identity or sequence. Uncertainties for such attributes may be expressed as probabilities.
NOTE 2 - Metrologically valid procedures for the production and certification of reference materials are given in, among others, ISO Guides 34 and 35.
NOTE 3 - ISO Guide 31 gives guidance on the contents of certificates.
NOTE 4 - VIM has an analogous definition (ISO/IEC Guide 99:2007, 5.14).

## Slide 7
An MKS PAMP Group Company
ISO Guides & Standards for
Reference Material

## Slide 8
ISO Guides relating to Reference Materials interlinking
An MKS PAMP Group Company
ISO Guide 34 competence of RM producer
ISO Guide 35 characterization and certification of RMs
ISO Guide 79 production of RMs for qualitative analysis
ISO Guide 80 in-house preparation of RMs for QC purposes
ISO Guide 33 uses of RMs includes calibration, method validation control charts etc.
ISO Guide 31 accompanying documentation certificates/labels/transport documents.
ISO Guide 30 definitions and terminology related to RMs.

## Slide 9
ISO Guide 34 and ISO/IEC 17025
An MKS PAMP Group Company
ISO Guide 34
Quality system requirements for RM production
Organizational requirements
Production control
ISO/IEC 17025
Specifies requirements for competence to carry out tests and calibrations using standard methods, on standard methods and laboratory developed methods.
Management requirement
Technical requirement
RM production control
Measurement system control
Significant overlap between ISO Guide 34,ISO/IEC 17025 and ISO 9001 (Quality management systems –Requirements)

## Slide 10
An MKS PAMP Group Company
Importance of
CRM

## Slide 11
Calibration
An MKS PAMP Group Company
use as a matrix matched calibrant (direct or via working standards) to ensure traceability of results to an external reference (the CRM)

## Slide 12
Validation
An MKS PAMP Group Company
check the measurement results in terms of validity:
is there any method specific bias ?
is there any systematic error ?
certified concentration

## Slide 13
Measurement Control
An MKS PAMP Group Company
Abuse of CRM: Misuse of CRM
Instead : use in-house materials or quality control materials
(i.e. of proven homogeneity and stability;
sometimes named in-house RMs or laboratory
RMs)
-3s
-2s
-1s
0
1s
2s
3s
18
17
16
15
14
13
12
11
10
9
8
7
6
5
4
3
2
1
(Upper) Control line
(Lower) Control line
(Upper) Warning line
(Lower) Warning line

## Slide 14
Use of CRMs
An MKS PAMP Group Company
How can CRMs help my measurements?
calibration (?)
validation (?)
measurement control (?)
X YES
X YES
X NO

## Slide 15
Importance of Measurement traceability
An MKS PAMP Group Company
Traceability - trace a product: from origin to manufacture & delivery/receipt.
Measurement traceability
In an absolute sense, the ‘true’ value can be defined only as being that value directly traceable to the base system of measurement (SI) or their derivatives – i.e: to national or international standards via an unbroken chain of comparisons
Provides
traceability
Transfers
accuracy
SI
unit
Certified
reference material
Laboratory
working standard
Analytical method

## Slide 16
An MKS PAMP Group Company
Standards for Precious Metal assaying

## Slide 17
Standards in use for Assaying: GOLD
An MKS PAMP Group Company
ISO 11426:2014(E)
Jewellery – Determination of gold in gold jewellery alloys – cupellation method (fire assay)
IS 1418 (2009)
Jewellery – determination of gold in gold bullion, alloys and jewellery/Artefacts – cupellation method (fire assay)
ASTM E1335-08
Determination of gold in bullion by fire assay cupellation analysis For Au in concentration range to 998.0 ‰ Classic fire assay method with gravimetry
AS 3515.2005
Determination of gold content -gravimetric(Fire Assay)method
ISO 15093:2014(E)
Determination of precious metals in 999‰ gold, platinum and palladium jewellery alloys —Difference method using ICP-OES

## Slide 18
Standards in use for Assaying: SILVER
An MKS PAMP Group Company
ISO 11427:2014(E)
Jewellery – Determination of silver in silver jewellery alloys – volumetric (potentiometric) method by KBr
IS 2113:2002
Assaying silver in silver and silver alloys – Methods: Gravimetric and volumetric(potentiometric)
ISO13756:2015
Jewellery – Determination of silver in silver jewellery alloys – volumetric (potentiometric) method by NaCl/KCl (optional for IS2113)
ISO 15096:2014
Determination of silver in 999‰ silver jewellery alloys —Difference method using ICP-OES

## Slide 19
An MKS PAMP Group Company
Comparison of Specifications
in Standards - GOLD

## Slide 20
An MKS PAMP Group Company
ASTM/BIS/ISO/AS Specifications & Limit Comparison for GOLD
 | IS 1418 | ISO 11426 |  | ASTM1335-08 |  | AS3515.1-205
Range ‰ | for 375-999 | for 333-990 | for ≥ 990 | for 200-990 | for ≥ 995 | For 300-999.5
Gold min ‰ | 999.9 | 999.9 | 999.99 | 999.9 | 999.99 | 999.96

## Slide 21
Specifications & Limit Comparison for GOLD at MMTC-PAMP :
An MKS PAMP Group Company
MMTC-PAMP/ASTM/BIS/ISO/AS Specifications & Limit(in ppm) Comparison for GOLD |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Impurities | Title |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 | Au-995‰ |  | Au-999‰ |  | Au-999.5‰ |  | Au-999.9‰ |  | Au-999.95‰ |  | Au-999.9‰ |  |  | Au-999.99‰ |  |  | 
 | MMTC-PAMP | ASTM | MPIPL | ASTM | MMTC-PAMP | ASTM B562-95 | MMTC-PAMP | ASTM B562-95 | MMTC-PAMP | ASTM B562-95 | IS 1418 | ISO 11426 |  | ASTM1335-08 |  | AS3515.1-205 | 
 |  |  |  |  |  |  |  |  |  |  |  | for 333-990 | for ≥ 990 | for 200-990 | for ≥ 995 |  | 
Gold min ‰ | 995.05 | Not specified | 999.05 | Not specified |  | 999.5 |  | 99.99 |  | 999.95 | 999.9 | 999.9 | 999.99 | 999.9 | 999.99 | 999.96 | 
Ag |  |  |  |  |  | 350 | 50 | 90 |  | 10 |  |  |  |  |  |  | 
Pt | addition ≤ 80 |  | 30 |  |  |  | 30 |  |  |  |  |  |  |  |  |  | 
Pd |  |  | 40 |  |  | 200 | 40 | 50 |  | 10 |  |  |  |  |  |  | 
Rh |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Ru |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
As |  |  |  |  |  |  | 30 | 30 |  |  |  |  |  |  |  |  | 
Bi |  |  |  |  |  |  | 20 | 20 |  | 10 |  |  |  |  |  |  | 
Cr |  |  |  |  |  |  | 3 | 3 |  | 3 |  |  |  |  |  |  | 
Cu | 500 |  |  |  |  | 200 | 20 | 50 |  | 10 |  |  |  |  |  |  | 
Fe | 50 |  |  |  |  |  | 20 |  |  | 10 |  |  |  |  |  |  | 
Hg |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Ir |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Mg |  |  |  |  |  |  | 30 | 30 |  | 10 |  |  |  |  |  |  | 
Mn |  |  |  |  |  |  | 3 | 3 |  | 3 |  |  |  |  |  |  | 
Ni |  |  |  |  |  |  | 3 | 3 |  |  |  |  |  |  |  |  | 
Os |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Pb | 30 |  |  |  |  | 50 | 10 | 20 |  | 10 |  |  |  |  |  |  | 
Sb |  |  |  |  |  |  | 10 |  |  |  |  |  |  |  |  |  | 
Se |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Si |  |  |  |  |  |  | 50 | 50 |  | 10 |  |  |  |  |  |  | 
Sn | 30 |  |  |  |  |  | 10 | 10 |  | 10 |  |  |  |  |  |  | 
Te |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 

## Slide 22
An MKS PAMP Group Company
Comparison of Specifications
in Standards - SILVER

## Slide 23
An MKS PAMP Group Company
ASTM/IS/ISO/AS Specifications & Limit Comparison for SILVER
Specifications | IS 1418 | IS2113 | ISO 11426 | ISO 11427 | ASTM1335-08 | AS3515.1-205
Silver min ‰ | 999 free from Au and PGM | 999 for gravimetric and 999.9 for Volumetric | 999 free from Au and PGM | 999.9 | 999 and 5 ppm gold maximum | 999 and less than 5 ppm gold

## Slide 24
Specifications & Limit Comparison for SILVER at MMTC-PAMP :
An MKS PAMP Group Company
MPIPL/ASTM/BIS/ISO/AS Specifications & Limit(IN PPM) Comparison for SILVER |  |  |  |  |  |  |  |  |  |  |  | 
Impurities | Title |  |  |  |  |  |  |  |  |  |  | 
 | Ag-999‰ |  | Ag-999.5‰ |  | Ag-999.9‰ |  | Ag-999‰ |  |  |  | Ag-999‰ | 
 | MMTC- PAMP | ASTM B413-97a | MMTC-PAMP | ASTM B413-97a | MMTC- PAMP | ASTM B413-97a | IS 1418 | IS2113 | ISO 11426 | ISO 11427 | ASTM 1335-08 | AS3515.1-205
Silver min ‰ | 999.05 | 999.0 | Not specified | 999.5 |  | 999.9 | 999 | 999/999.9 | 999 | 999.9 | 999 | 999
Au | 30 |  |  |  | 30 |  | 0 |  | 0 |  | < 5 | < 5
Pt |  |  |  |  |  |  | 0 |  | 0 |  |  | 
Pd |  |  |  |  | 10 | 10 | 0 |  | 0 |  |  | 
Rh |  |  |  |  |  |  | 0 |  | 0 |  |  | 
Ru |  |  |  |  |  |  | 0 |  | 0 |  |  | 
Bi | 20 | 10 |  | 10 | 8 | 5 |  |  |  |  |  | 
Cu | 500 | 800 |  | 400 | 30 | 100 |  |  |  |  |  | 
Fe | 20 | 20 |  | 20 | 10 | 10 |  |  |  |  |  | 
Ir |  |  |  |  |  |  | 0 |  | 0 |  |  | 
Os |  |  |  |  |  |  | 0 |  | 0 |  |  | 
Pb | 250 | 250 |  | 150 | 10 | 10 |  |  |  |  |  | 
Sb |  |  |  |  | 10 |  |  |  |  |  |  | 
Se |  |  |  |  | 5 | 5 |  |  |  |  |  | 
Te |  |  |  |  | 5 | 5 |  |  |  |  |  | 
 | PGM group metals |  |  |  |  |  |  |  |  |  |  | 

## Slide 25
An MKS PAMP Group Company
What MMTC-PAMP can provide

## Slide 26
Gold-Silver for Reference Material
An MKS PAMP Group Company
GOLD:
Fine Gold of 999.9 fineness complying with International specifications.
This can be provided in wire of 0.9mm Diameter
Fine Gold of 999.99 fineness complying with International specifications.
This can be provided in wire of 0.9mm Diameter.
SILVER:
Fine Silver of 999 fineness complying with International specifications.
This can be provided in grain form or wire form.
Fine Silver of 999.9 & 999.99 fineness complying with International specifications.
This can be provided in grain forms or wire form.

## Slide 27
Any questions
An MKS PAMP Group Company
Any questions ???????

## Slide 28
THANK YOU